## Interface options
Use these and ONLY these Interface options:
Interface/function | Address/parameter | Command | STATUS
GPIB | [GPIB address] |  | TESTED
RS232 |  |  | ?
VISA | [VISA ID] |  | TESTED
TCP/IP | [IP:port] |  | TESTED |  | Tested with CLIMATS chamber with some Spirale3 specific configuration
BURSTER_4462 | [VISA ID] |  | TESTED
|  |  | TESTED |  | Empty "Interface" lines should be completely ignored, including other parameters in this line. Therefore, you can place there your comments.
STOP | [no parameter] |  | TESTED |  | Not really an interface, but stops the execution loop. Other parameters are ignored. If no STOP is set, the script will loop infinitely until it will be stopped manually.
FILE | COMMENT | Comment text | TESTED |  | Write a comment to file. Wrire the text in the COMMAND column.
FILE | ENDL |  | TESTED |  | Write a ENDLine character to the file.
FILE | [other FILE parameters] |  | NOT TESTED
PAUSE | [no parameter] | Dialog text | TESTED |  | Waits for pressing button before continue. Write popup description in COMMAND column. Button options are "continue" and "exit"
CALL_SUBSCRIPT | [NUMBER or NAME] | number or name | TESTED |  | Call a subscript with a specified name or number (using of numbers is not reccomended, because it takes care of the sheet order in excel)
CALL_SUBSCRIPT | NAME | name;;<param1>==val1;;<param2>==val2 | TESTED |  | Call a subscript like a function (with parameters to replace)
EXEC | [Wait?;;Minimize?] | .exe call | TESTED |  | call an executable without or with parameters
BMD_DIAG_TIMER | [VISA ID] | name:time[s] | TESTED |  | get values over LIN (use ":" as parameter separator)
GPT9000 | [VISA ID] | ACW_min_step_max_time | TESTED
WAIT | NA | NA | TESTED |  | Waits the amount of seconds
others will not be recognized properly.
Remember that VISA needs additional configuration in MAX.
Remember that you have to check the serial configuration if serial port is used
Example for using GPIB
Interface | Address | Read response? | Save response? | Wait after [s] | COMMAND
GPIB | 9 | 0 | 1 | 30 | ROUT:CLOS (@301) |  |  | Obviously, a save request will be ignored, if there is no response.
| Interface | GPIB | The spelling must be identical.
| Address | GPIB Address according the setting in the instrument. Follow the user manual of the instrument in order to enable the GPIB communication and set the address. Address can be between 1 and 16.
| Read response? | Can be 0 or 1 | 0 is for NO and 1 for YES. This setting is depending on the protocol only and must correspond to it. In case of mismatch, you will get an error.
| Save response? | Can be 0 or 1
| Wait after [s] | Can be any; defines the number of seconds to wait before the next line is executed
| COMMAND | Command; refer to the programming manual of the instrument to find out the proper command. You can find some examples beside.
CAUTION: You can not have multiple GPIB interfaces in the system. Use VISA in case of problems.
Example for using RS232
Interface | Address | Read response? | Save response? | Wait after [s] | COMMAND
RS232
| Interface | RS232 | The spelling must be identical.
Address
| Read response? | Can be 0 or 1 | 0 is for NO and 1 for YES. This setting is depending on the protocol only and must correspond to it. In case of mismatch, you will get an error.
| Save response? | Can be 0 or 1
| Wait after [s] | Can be any; defines the number of seconds to wait before the next line is executed
| COMMAND | Command; refer to the programming manual of the instrument to find out the proper command. You can find some examples beside.
Example for using VISA
Interface | Address | Read response? | Save response? | Wait after [s] | COMMAND
VISA
| Interface | VISA | The spelling must be identical. |  | VISA is an alternative to use low-level GPIB or serial communication
| Address | ID assigned in the front panel; in order to use VISA, you have to select the visa instrument on the front panel and assign a reference ID - use this ID here.
| Read response? | Can be 0 or 1 | 0 is for NO and 1 for YES. This setting is depending on the protocol only and must correspond to it. In case of mismatch, you will get an error.
| Save response? | Can be 0 or 1
| Wait after [s] | Can be any; defines the number of seconds to wait before the next line is executed
| COMMAND | Command; refer to the programming manual of the instrument to find out the proper command. You can find some examples beside.
Multiple commands can be separated using ";", if the instrument accepts that
Example for using TCP/IP
Interface | Address | Read response? | Save response? | Wait after [s] | COMMAND
TCP/IP | 192.168.141.107:6040 | 1 | 0 | 10 | Marche_arret = 1 |  | //enable CLIMATS (CC-25)
| Interface | TCP/IP | The spelling must be identical.
| Address | The format is IP:PORT |  | Refer to the user manual to enable the ethernet communication and set the IP address.
| Read response? | Can be 0 or 1 | 0 is for NO and 1 for YES. This setting is depending on the protocol only and must correspond to it. In case of mismatch, you will get an error.
| Save response? | Can be 0 or 1
| Wait after [s] | Can be any; defines the number of seconds to wait before the next line is executed
| COMMAND | Command; refer to the programming manual of the instrument to find out the proper command. You can find some examples beside.
Example using BURSTER_4462
Interface | Address/ID | Read response? | Save response? | Wait after [s] | COMMAND
BURSTER_4462 | Burster | 1 | 1 | 0 | SOUR:VOLT 0.12
| Interface | BURSTER_4462 | The spelling must be identical.
| Address | ID assigned in the front panel; in order to use VISA, you have to select the visa instrument on the front panel and assign a reference ID - use this ID here.
| Read response? | This parameter is ignored because for Burster Digistant 4462 response always must be read and this is already implemented in the sub vi function.
| Save response? | Can be 0 or 1
| Wait after [s] | Can be any; defines the number of seconds to wait before the next line is executed
| COMMAND | Command; refer to the programming manual of the instrument to find out the proper command. You can find some examples beside.
Example using CALL_SUBSCRIPT
CALL_SUBSCRIPT is a pseudo interface which allows to extract repeated and long parts of script and include them with only 1 line to the main script.
Interface | Address/ID | Read response? | Save response? | Wait after [s] | COMMAND
CALL_SUBSCRIPT | NUMBER or NAME | 0 | 0 | 0 | number or name
| Interface | CALL_SUBSCRIPT | The spelling must be identical.
| Address | "NUMBER" for excel sheet number; "NAME" for excel sheet name
| Read response? | NA
| Save response? | NA
| Wait after [s] | NA
| COMMAND | Excel sheet number, where first is "0" - If Address == NUMBER; Excel sheet name if Address == NAME
Note: A "STOP" command is removed from the subscript
Call subscript with replace
CALL_SUBSCRIPT is a pseudo interface which allows to extract repeated and long parts of script and include them with only 1 line to the main script.
Interface | Address/ID | Read response? | Save response? | Wait after [s] | COMMAND
CALL_SUBSCRIPT | NUMBER or NAME | 0 | 0 | 0 | name;;<param1>==val1;;<param2>==val2
| Interface | CALL_SUBSCRIPT | The spelling must be identical.
| Address | "NUMBER" for excel sheet number; "NAME" for excel sheet name
| Read response? | NA
| Save response? | NA
| Wait after [s] | NA
| COMMAND | Excel sheet number, where first is "0" - If Address == NUMBER; Excel sheet name if Address == NAME
;; is defauld sheet-from parameters and parameters delimiter
== is default replacement operator
Replaced fields are by default : |  | Wait after [s], COMMAND, COMMENT
A subscript can now contain another subscript and so on.
If main script is loaded from txt, subscript names will be searched in txt files (do not include .txt extension in command field
Note: A "STOP" command is removed from the subscript
EXEC program call
Call an executable
Interface | Address/ID | Read response? | Save response? | Wait after [s] | COMMAND
EXEC | 1;;0 | 1 | 0 | 5 | USBlini_lib_s54.exe -r 1 -t 15
| Interface | EXEC | The spelling must be identical.
| Address | <par1> | Wait until completion? |  | 0 or 1
|  | ;; | parameter separator
|  | <par2> | Run minimized? |  | 0 or 1 | Hides the cmd window
| Read response? | Parses the output from cmd
Save response?
| Wait after [s] | Waits the amount of seconds. Depending of the parameter in the adress field, the wait begins after completion of the executable or after it has been runned
| COMMAND | system command, example:
|  |  | USBlini_lib_s54.exe |  |  | Call without parameters
|  |  | USBlini_lib_s54.exe -r 1 -t 15 |  |  | Call with some parameters
|  |  | D:\!D\04_Programiki\SEx\builds\SEx_master\SEx 5.10.1.0\USBlini_lib_s54.exe | [5 empty] | Call without parameters with a specified location
|  |  | D:\!D\04_Programiki\SEx\builds\SEx_master\SEx 5.10.1.0\USBlini_lib_s54.exe -r 1 -t 15 | [5 empty] | Call with some parameters with a specified location
BMD data logging over diagnostic frames
Use LIN interface from "6-fach tester"/"functional tester" to logg I, U12, U24 and T
Interface | Address/ID | Read response? | Save response? | Wait after [s] | COMMAND
BMD_DIAG_TIMER | LIN_DUT1 | 1 | 1 | 0 | Unom_meas_Delta@0A:30
| Interface | BMD_DIAG_TIMER |  | Refers to the LIN interface from "6-fach BMD functional tester"
| Address | LIN_DUT1 |  | VISA alias or COM port
Read response?
| Save response? |  |  | if save is selected, last result is saved
Wait after [s]
| COMMAND | Unom_meas_Delta@0A:30
|  |  | Unom_meas_Delta@0A |  | comment for the log file
|  |  | : |  | command separator
|  |  | 30 |  | time in seconds for logging
Perform ACW up-to-fail test on GW INSTEK GPT9000
Interface | Address/ID | Read response? | Save response? | Wait after [s] | COMMAND
GPT9000_ACW_UTF | COM4 | 1 | 1 | 0 | ACW_1,0_0,1_5,0_60
| Interface | GPT9000 |  | Function reference
| Address | [VISA ID] |  | VISA Alias
| Read response? | 1 |  | you shall read the response
| Save response? | 1 |  | you shall save the response
Wait after [s]
| COMMAND | ACW_1,0_0,1_5,0_60 |  | exemplary command
|  |  | ACW | refers to device function; currently not changable
|  |  | _ | separator
|  |  | 1,0 | start voltage in [kV]
|  |  | _ | separator
|  |  | 0,1 | step voltage in [kV]
|  |  | _ | separator
|  |  | 5,0 | end voltage in [kV]
|  |  | _ | separator
|  |  | 60 | step time in [s]

## Buttons
Log commands?
Enables you to log commands line-by-line.
This works only for request response instrument commands.
Log comments?
Enables you to log comments.
This works only for request response instrument commands.
Change ENDL to TAB mode?
Turn ON if you want to replace the default ENDL file character to TAB after logging the current response.
Reverse TIME-NAME for logfile?
Reverse the order for TIME and NAME of the result filename.

## Commands library
Equipment | COMMANDS | Comment
or function
AGILENT 34970A | ROUT:CLOS (@301)
CONF:VOLT:DC 1.0,0.00001,(@120)
SENS:VOLT:DC:NPLC 200,(@120)
READ?
CONF:VOLT:DC 0.1,0.000001,(@219)
SENS:VOLT:DC:NPLC 20,(@219)
READ?
CONF:VOLT:DC 0.1,0.000001,(@119)
SENS:VOLT:DC:NPLC 200,(@119)
READ?
ROUT:OPEN (@301)
SENS:VOLT:DC:NPLC? (@119)
SENS:VOLT:DC:NPLC? (@120)
SENS:VOLT:DC:NPLC? (@219)
MEAS:VOLT:DC? 100,0.001,(@105)
MEAS:VOLT:DC? 10,0.00001,(@106)
MEAS:VOLT:DC? 100,0.001,(@101)
MEAS:TEMP? TC,K,(@103)
MEAS:VOLT:DC? 100,0.001,(@102)
MEAS:TEMP? TC,K,(@104)
| MEAS:FRES? (@110) | 4-wire resistance
| MEAS:RES? (@110) | 2-wire resistance
CONF:VOLT:DC (@103)
ROUT:MON:CHAN (@103)
ROUT:MON:STAT ON
| CONF:FRES 1000,1,(@301,302) | config 4-wire resistance with 1R resolution
| SYST:CPON 200 | card reset, slot 200
Burster Digistant 4462 | SOUR:VOLT 0.5
SOUR:VOLT?
| SYST:ERR? | Reads the last error from stack.
SOURce:VOLTage:RANGe:AUTO?
SOURce:VOLTage:RANGe:AUTO ON
SOURce:VOLTage:RANGe:AUTO OFF
SOURce:CURRent:DELTa?
SOURce:VOLTage:DELTa?
SOURce:VOLTage:ALC:SOURce?
| SOURce:VOLTage:ALC:SOURce INT | Voltage sense INTERNAL
| SOURce:VOLTage:ALC:SOURce EXT | Voltage sense EXTERNAL
SOURce:VOLTage:PROTection:LEVel?
SOURce:CURRent:PROTection:LEVel?
SOURce:MODE?
| SOURce:MODE FIXed | Normaler Gebermode
| SOURce:MODE SWEep | Rampe1 (gleicher Gebermode, gleiche ∆-Werte und gleiche ∆-Zeit)
| SOURce:MODE LIST | Rampe2 (verschiedene Gebermodes, verschiedene ∆-Werte, verschiedene ∆-Zeiten
| SOURce:SWEep:WAVeform? | Die Ablaufform von Rampe1 wird abgelesen.
| SOURce:SWEep:WAVeform TRIangle | Ablauf in Form eines Dreiecks
| SOURce:SWEep:WAVeform SAWTooth | Ablauf in Form eines Sägezahns
| SOURce:SWEep:COUNt? | Anzahl der eingestellten Durchläufe von Rampe1: 0 bis 99
| SOURce:SWEep:COUNt <parameter> | Anzahl der Durchläufe: 0 bis 99 (0 bedeutet Dauerlauf).
SV?
| SV -5.5 | set voltage
SC?
| ST 100C | set temperature simulation (deg. C)
ST?
GEN 40-250 | SOUR:VOLT XX.XX
SOUR:CURR XX.XX
OUTP:STAT OFF
OUTP:STAT ON
MEAS:CURR?
1800A Load | MEAS:CURR? | //ELoadMeasuredCurrent, DischargeCurrent_ELoad, ELoadCurrentSetValue
MODE:CURR
CURR XXX
INP OFF
INP ON
| SYST:CONT EXT | //set eload to ext. contr.
| CURR:RANG 1800 | //set range to 1800 A
INP ON
| CURR:RANG 1 | //set range to 1 A
CURR 0
INP OFF
*RST
Agilent 33220A Generator | DISP:TEXT 'Real Crank Start
| FUNC:USER CRANK_1 | //select crank pulse as waveform
| FUNC:SHAP USER | //select crank pulse as waveform next command
| FREQ 0.0312 | //set freq 31,2 mHz
| VOLT:AMPL 1.85 | //set Vpp
| VOLT:OFFS 0 | //set offset to 0
| BURS:MODE TRIG | //set triggemode
| BURS:NCYC 1 | //set burst cycle to 1
| TRIG:SOUR BUS | //select trigger source to GPIB
| BURS:STAT ON | //enable trigger mode
OUTP ON
| *TRG | //trigger over GPIB
OUTP OFF
Agilent 34410A | CONF:VOLT:DC 100, 0.3
MEAS:VOLT?
CONF:VOLT:DC 10, 0.3
MEAS:VOLT?
MEGA HV Relay module | K04=1 | //enable K04 relay
GPT9000 | ACW_1,0_0,1_5,0_60 | //executes a "up to fail" ACW test starting from 1,0 kV up to 5,0 kV with 0,1 kV steps, for 60 seconds each step.
//other parameters can not be changed at the command level
CLIMATS CC | ConsCuve = 25 | set temperature
| ConsHum = 0 | set humidity
| Marche_arret = 1 | enable chamber, wait 3240s/54min
| ?TCuve | ask for temperature
| ?HumCuve | ask for humidity
Hameg HMP2030 | *IDN?
| OUTPut:GENeral ON | enable general output
OUTPut:GENeral OFF
INSTrument OUT3
OUTPut:SELect OFF
INSTrument OUT2
OUTPut:SELect OFF
INSTrument OUT1
APPLy 10.0, 0.2
OUTPut:SELect ON
| INSTrument OUT2 | always set a channel first
| APPLy 9.0, 0.3 | set voltage and current at the same time: APPLy <volt>, <curr>
| OUTPut:SELect ON | enable only this channel (if output:general is off, output is still off)
INSTrument OUT3
APPLy?
VOLTage?
CURRent?
EA PS | *IDN?
| SYSTem:LOCK ON | enable remote control
SOURce:VOLTage 0.0; SOUR:CURR 0.0
SOURce:POWer MAX
OUTPut ON
SOURce:VOLTage 10.0; SOUR:CURR 5.0
OUTPut OFF
OUTPut ON
MEASure:VOLTage?
MEASure:CURRent?
MEASure:POWer?
| SYSTem:LOCK OFF | disable remote control

## exemplary script
Interface | Address/ID | Read response? | Save response? | Wait after [s] | COMMAND | COMMENT
GPIB | 6 | 1 | 0 | 0 | OUTP:STAT OFF
VISA | MEGA | 1 | 1 | 0 | K37=0
VISA | MEGA | 1 | 1 | 0 | K35=1
GPIB | 6 | 1 | 0 | 0 | SOUR:VOLT 8.00
GPIB | 6 | 1 | 0 | 0 | SOUR:CURR 0.00
GPIB | 6 | 1 | 1 | 0 | MEAS:VOLT?
GPIB | 6 | 1 | 0 | 0 | OUTP:STAT ON
GPIB | 6 | 1 | 0 | 60 | SOUR:CURR 1.00
GPIB | 6 | 1 | 0 | 60 | SOUR:CURR 5.00
GPIB | 6 | 1 | 0 | 60 | SOUR:CURR 10.00
GPIB | 6 | 1 | 0 | 60 | SOUR:CURR 15.00
GPIB | 6 | 1 | 0 | 60 | SOUR:CURR 20.00
GPIB | 6 | 1 | 0 | 60 | SOUR:CURR 50.00
GPIB | 6 | 1 | 0 | 60 | SOUR:CURR 75.00
GPIB | 6 | 1 | 0 | 60 | SOUR:CURR 100.00
GPIB | 6 | 1 | 0 | 15 | SOUR:CURR 125.00
GPIB | 6 | 1 | 0 | 15 | SOUR:CURR 150.00
GPIB | 6 | 1 | 0 | 15 | SOUR:CURR 200.00
GPIB | 6 | 1 | 0 | 0 | SOUR:CURR 0.00
GPIB | 6 | 1 | 0 | 0 | OUTP:STAT OFF